## Tabelle1
OSX Conductivity V1.0 – Calibration – 04.08.2023

	Measure 0	Measure 1
0: Cond (µS/cm)
1: Temp (°C)	24
2: Voltage (V)
3: Cond (µS/cm, uncomp)	138	1332
4: Cond, Raw Signal	85605	527804
5: Ref Signal	369545	(only for K6 = 1.0)

	Standard	Scientific		Temp.
K0 – Cond Multi	0,00270014179136543	2,7001418E-03		442199	(DX)
K1 – Cond Offset	93,1456380498373	9,3145638E+01		1194	(DY)
K2 – Temp Multi	(unchanged, 1.0)
K3 – Temp Offset	-1
K4 – Temp TK
K5 – Gain
K6 – Ref Multi	2,70603038872127E-06	2,7060304E-06	(only for K6 = 1.0)
K7 – Ref Offset	(unchanged, 0.0)

Conductivity Calibration:
- Make 2 Measures ('e9') in Calibration Solutions (e.g. 147 µS/cm and 1413 µS/cm)
- The Calibration Solutions have small Temperature Dependency (e.g. 138.4 µS/cm @ 22°C and 1332 µS/cm @ 22°C). These are the uncompensated Values:
- Enter the uncompensated Conductivity Values (as Setpoint) and Raw Signals for Measure 0 and 1
- Set the calculated K0,K1 with 'z?K0=xxx!' or 'z?K1=xxx!' for the new Conductivity Calibration Coefficients

Reference Signal Calibration (optional! Only as info)
- Optionally set K6 to 1.0 with ‘z?K6=1!‘
- Measure with ‘e9‘

Temperature Offset Calibration
- Enter measured Temperature (e.g. 24 )
- Set the calculated Offset with 'z?k3=xxxx!'

Don’t forget to save the changed Coefficients with ‘z?Xwrite!‘